# Firmware Version Update Procedures

Review the following precautions before performing a firmware version update.

Failure to follow these precautions could result in the destruction of the internal flash memory and prevent the RM1100 from starting properly.

＜Precautions＞

①　Always plug in the AC adapter before performing a firmware update.

Do not cut the power while the device is updating.

(Never attempt a firmware update while the device is running on battery power)

② Do not remove the SD card while the device is updating.

③ Do not include both ”Main FW” and “Initialization Data”files on the same SD used for the firmware update.

Always perform the data with each file on a separate card.

④ Do not use non-designated SD cards.

＜Necessary materials and software＞

① SD Card x2 （One each for the most recent version of the following files）

・Main firmware update file　“RM1100PG.bin”

・Initialization data file “RM1100FI.txt”

＜Main firmware version update procedures＞

① Power the device off and insert the SD card containing the main firmware update file (RM1100PG.bin）.

② Power on the device.

③ The firmware update screen (blue) will display. Wait approximately three minutes for the screen to turn black.

Caution： Updating may take a while. Do not touch the device during this time.

Approx. 3 mins

④ Turn off the device and remove the SD card.

＜Device initialization＞

① Power the device off and insert the SD card containing the initialization data file.

② Power on the device.

③ Once the application screen appears, remove the SD card. (The power can be left on)

Caution： Do not touch the device until the application screen appears.

＜Confirm firmware update＞

By navigating through the following, “SYSTEM”→ “SERVICE” → “VERSION”, “MAIN PROGRAM” will display as below.

MAIN PROGRAM： V1.1Ｉ

Confirm the version of “MAIN PROGRAM”